## Before you begin
How to use this workbook
Clinical validation is carried out in a front-to-end fashion, i.e. covering ALL steps included in the test. Traditionally, this includes all steps from DNA to final result (clinical report) but today, parts of the NGS process (such as sequencing, bioinformatics or clinical interpretation) are offered separately by other clinical entities. The validation requirements for these partial services are different from those applied to a full service laboratory and are described separately below.
Please note: Validations for supplemental assays may not be covered by the illustrated NGS validation setup in subsequent tabs but can consider using a similar framework.
1: Design your validation study
The following hybrid validation design is recommended for LDTs (laboratory developed tests)
A) Baseline analytical performace
Establish general analytic performance by variant type that must be detected by the test (i.e. SNVs, indels, CNV, other) and each specimen type the test must accommodate (e.g. whole blood, saliva, amniotic fluid, etc.)
For variants other than SNVs, establish performance for different categories (e.g. different size ranges for indels or CNVs; see "GENERAL RESOURCES" tab)
Maximize # of variants tested for increased statistical confidence
Note: This component is "methods-based" (Aziz et al. 2015, [checklist codes]), i.e. performance data from multiple tests (e.g. multiple gene panels) can be added to enhance statistical power provided that the test protocols are identical
Commonly used samples for this type of validation analysis are highly characterized samples, such as those characterized by the Genome in a Bottle Consortium (http://jimb.stanford.edu/giab/; Zook 2016 PMID 27271295) or CDC Genetic Testing Reference Materials Coordination Program (GeT-RM; https://wwwn.cdc.gov/clia/Resources/GetRM/).
B) Additional studies as needed. Examples include but are not limited to:
Medically critical gene with high prevalence of a variant type where the base-line performance does NOT provide sufficient statistical confidence (e.g. the number of variants for a certain variant type that can be derived from the reference sample/s is too low). Example: CNV detection for DMD (Duchenne Muscular Dystrophy)
Common pathogenic variant/s exist and base-line performance does NOT provide sufficient confidence for the type of variant. Example: the common, recurring 3 base deletion in the CFTR gene (c.1521_1523delCTT; p.F508del)
Gene or clinically relevant variant falls within a difficult to sequence region, which may have different analytical performance metrics than general regions analyzed in methods-based validation. Example: Genes with regions of high homology to other regions of the genome (e.g.PMS2, STRC, Mandelker 2016)
2. Start on Tab A, Performance Metrics
Review performance metrics that need to be recorded
The "PERFORMANCE METRICS" tab provides a comprehensive summary of analytical performance metrics that are typically assessed during validation. Note that terminology and definitions is not entirely harmonized across different publications/recommendations that have been issued to date (list of such publications is provided on the "GENERAL_RESOURCES" tab).
NGS testing that is designed to detect variants represents a qualitative test where the signal obtained is translated to a limited number or pre-defined outcomes (homozygous reference, heterozygous reference/variant, homozygous variant).
While there are many acceptable approaches to satisfy the commonly used analytical performance metrics, keep in mind that the following are the core principles that are addressed:
1. How well does the assay detect variants that are known to be present?
2. How high is the likelihood for detecting variants that are not present (false positive results)?
3. How "robust" is the assay when certain conditions are altered (e.g. a different sequencing machine is used)?
3. Continue to Tabs C and D (Example Sample List and Example Validation Grid)
Assemble your validation samples and design validation runs
Please note that these tabs illustrate ONE possible approach that illustrates how a validation can be carried out.
4. Continue to Tabs E and F (Example Performance Metrics and Assay Limitations)
After analyzing the validation runs and calculating the analytical performance metrics (see "EXAMPLE PERFORMANCE METRICS" tab), review the achieved performance and ensure that it meets requirments for the intended use. If those requriements are NOT met, go back to "TEST OPTIMIZATION" worksheet. Otherwise, the test can be launched. It is advisable to summarize limitations (see "ASSAY LIMITATIONS" tab for a representative example)

## A_Performance_Metrics
PERFORMANCE METRIC | DEFINITION | DEFINITION (DETAIL) | ADDITIONAL COMMENTS | EXAMPLE APPROACHES
A) Variant detection
Analytical sensitivity (MOL.31360) | TP/[TP+FN] | The fraction of variants known to exist by a designated gold standard that are correctly identified as such by the new test | 1) Should be calculated separately for all types of variants relevant to the test/disease tested (confidence intervals should be calculated) 2) Values for separate tests (i.e. different groups of genes) can be added to derive a cumulative value provided that assay methodology is the same | - Use well characterized reference samples (e.g. NA12878, Coriell samples) - Use clinical specimens that have been tested by a clinically validated assay - Alternative reference materials (e.g. for rare but essential variants) such as engineered cell lines, linearized plasmids or other artificial constructs diluted in genomic DNA - In silico mutagenized files may be used to supplement patient samples and physical specimens, and these can be especially helpful in generating difficult to source variants (e.g., indels of varying sizes)
Analytical specificity (MOL.31375) | TN/[TN+FP] | The test's ability to obtain negative results in concordance with negative results obtained by the reference method (Ability to accurately call WT (reference) bases
Technical positive predictive value | TP/[TP+FP] | Likelihood that a variant call is a true positive
Accuracy (MOL.31130) | (TP+TN)/[TP+FP+TN+FN] | Accuracy can be equated to sensitivity + specificity. Note that this metric is not commonly used (Jennings 2009 and Mattocks 2010)
B) Performance across altered conditions and equipment
Precision (MOL.31145) |  | Closeness of agreement between independent test results obtained under stipulated conditions (a measure of random error) | Precision is usually divided into two sub-categories: 1. repeatability (within-run precision) 2. reproducibility (conditions are varied) | Performance can be measured using a variety of QC metrics established by the laboratory. Examples can include (not limited to): fragment size and distribution, library concentration, average coverage, % reads on target. If using reference specimen such as NA12878 then, analytical sensitivity and specificity can also be used as measures (refer to Example Validation Grid)
Repeatability (within-run precision) |  | Degree to which the same sequence is derived when sequencing a reference sample multiple times, under the same conditions
Reproducibility (conditions altered) |  | Closeness of agreement between repeated tests for the same sample under changed conditions (instrument, laboratory, operator, sample concentration, sample source, reagent lot, run, day, time of day)
Robustness |  | Test precision given small, deliberate changes in test conditions (e.g. variations in storage conditions
C) Other metrics
Limit of Detection (MOL.36118) |  | Lowest amount/concentration of analyte that gives an informative result at a given level of confidence
Reportable range (MOL.31245) |  | CAP checklist: The reportable range encompasses the full range of reported values. For qualitative tests that would include all reportable outcomes (e.g. homozygous wild type, heterozygous or homozygous variant)
Reference interval (MOL.31255) |  | The reference interval is the range of results expected in the normal population | For NGS, this refers to the normal variation of sequence within the population that the assay is designed to detect
GLOSSARY | REFERENCES
TP = True Positive | 1. Gargis 2012. PMID 23138292 (Assuring the quality of next-generation sequencing in clinical laboratory practice)
FP = False Positive | 2. Rehm 2013. PMID 23887774 (ACMG clinical laboratory standards for next-generation sequencing.)
TN = True Negative | 3. Mattocks 2010. PMID 20664632 (A standardized framework for the validation and verification of clinical molecular genetic tests)
FN = False Negative | 4. Jennings 2009. PMID 19415949 (Recommended principles and practices for validating clinical molecular pathology tests)
5. CSLI MM17
ADDITIONAL METRICS

## B_General_Resources
Calculating analytical performance
Reference Sample
|  | + | - |  | Analytical Sensitivity | TP/[TP+FN]
New Assay | + | TP | FP | TP+FP | Analytical Specificity | TN/[TN+FP]
| - | FN | TN | FN+TN | False Negative Rate | 1 - Sensitivity
|  | TP+FN | FP+TN |  | False Positive Rate | 1 - Specificity
Confidence interval calculator
http://www.pedro.org.au/english/downloads/confidence-interval-calculator/
Reference Materials
National Institute of Standards and Technology (NA12878, RM8391, RM8392, RM8393) | [6 empty] | https://www.nist.gov/programs-projects/genome-bottle
Genetic Testing Reference Materials Coordination Program (GeT-RM) | [6 empty] | https://wwwn.cdc.gov/clia/Resources/GETRM/
EMQN | [6 empty] | https://www.emqn.org/
Additional Publications
Kalman et al. Ann Lab Med. 2016 Nov;36(6):513-520.
Zook et al. Scientific Data 3, Article number: 160025 (2016)
Aziz 2015. Arch Pathol Lab Med.139:481 (College of American Pathologists' laboratory standards for next-generation sequencing clinical tests)

## C_Example_Sample_List
SAMPLE ID | RUN ID | WITHIN-RUN ID | DESCRIPTION | PURPOSE | PURPOSE (DETAIL) | SPECIFIC VARIANT/S OF INTEREST | VARIANT TYPE | MEASURED METRIC
This example reflects a validation scheme designed for the MiSeq: 10 samples/lane, 3 replicate runs
VAL-1 | run-1 | sample-1 | NA12878 | base-line validation | N/A | N/A | N/A | Intra-run variability | Concept: Base-line validation satisfied by NA12878 (or other well characterized reference sample/s) Fill run up with other reference materials depending on gene/disease specific needs
VAL-2 | run-1 | sample-2 | NA12878 | base-line validation | N/A | N/A | N/A | Intra-run variability
VAL-3 | run-1 | sample-3 | NA12878 | base-line validation | N/A | N/A | N/A | Intra-run variability
VAL-4 | run-1 | sample-4 | Anonymized patient specimen | base-line validation | variant type prevalent in gene | list variant here | IN/DEL | sensitivity
VAL-5 | run-1 | sample-5 | Anonymized patient specimen | base-line validation | variant type prevalent in gene | list variant here | CNV | sensitivity
VAL-6 | run-1 | sample-6 | Anonymized patient specimen | gene-specific validation | mosaicism (variant at 20%) | list variant here | SNV | limit of detection
VAL-7 | run-1 | sample-7 | Anonymized patient specimen
VAL-8 | run-1 | sample-8 | Anonymized patient specimen
VAL-9 | run-1 | sample-9 | Anonymized patient specimen
VAL-10 | run-1 | sample-10 | Anonymized patient specimen
VAL-11 | run-2 | sample-1 | NA12878 | base-line validation | N/A | N/A | N/A | Inter-run variability
VAL-12 |  | sample-2 | Anonymized patient specimen
VAL-13 |  | sample-3 | Anonymized patient specimen
VAL-14 |  | sample-4 | Anonymized patient specimen
VAL-15 |  | sample-5 | Anonymized patient specimen
VAL-16 |  | sample-6 | Anonymized patient specimen
VAL-17 |  | sample-7 | Anonymized patient specimen
VAL-18 |  | sample-8 | Anonymized patient specimen
VAL-19 |  | sample-9 | Anonymized patient specimen
VAL-20 |  | sample-10 | Anonymized patient specimen
VAL-21 | run-3 | sample-1 | NA12878 | base-line validation | N/A | N/A | N/A | Inter-run variability
VAL-22 |  | sample-2 | Anonymized patient specimen
VAL-23 |  | sample-3 | Anonymized patient specimen
VAL-24 |  | sample-4 | Anonymized patient specimen
VAL-25 |  | sample-5 | Anonymized patient specimen
VAL-26 |  | sample-6 | Anonymized patient specimen
VAL-27 |  | sample-7 | Anonymized patient specimen
VAL-28 |  | sample-8 | Anonymized patient specimen
VAL-29 |  | sample-9 | Anonymized patient specimen
VAL-30 |  | sample-10 | Anonymized patient specimen

## D_Example_Validation_Grid
SAMPLE VALIDATION RUN DESIGNS
| Example 1: |  |  |  |  | Example 2:
| HiSeq 2500 (8 lanes) |  |  |  |  | Miseq (1 lane)
| 10 samples/capture, run 3 times (same run) + 2 times on 2 additional runs |  |  |  |  | 10 samples/capture
load same pool 3 times | RUN 1 - LANE 1 | LANE 2 | LANE 3 | LANE 4 | LANE 5 | LANE 6 | LANE 7 | ◄--- | INTRA-RUN Run same pool of 10 three times on same flow cell (assesses instrument variability) | ---► | RUN 1 - LANE 1
| NA12878 - A - barcode 1 | NA12878 - A - barcode 1 | NA12878 - A - barcode 1 | OPTIONAL LANES, MAY HAVE NEED TO VALIDATE ADDITIONAL SAMPLES 1. Samples with frequent pathogenic variants 2. Samples carrying variant type/s of interest (e.g. In/Dels, CNVs, structural variants) | [6 empty] | NA12878 - A - barcode 1 | NOTE Intra-run variability limited to NA12878
| NA12878 - A - barcode 2 | NA12878 - A - barcode 2 | NA12878 - A - barcode 2 | [7 empty] | NA12878 - A - barcode 2
| NA12878 - A - barcode 3 | NA12878 - A - barcode 3 | NA12878 - A - barcode 3 | [7 empty] | NA19240 - A
| NA19240 - A | NA19240 - A | NA19240 - A | [7 empty] | Sample 4
| Sample 5 | Sample 5 | Sample 5 | [7 empty] | Sample 5 | NOTE Samples 4-10 may differ between runs based on need to add disease specific samples
| Sample 6 | Sample 6 | Sample 6 | [7 empty] | Sample 6
| Sample 7 | Sample 7 | Sample 7 | [7 empty] | Sample 7
| Sample 8 | Sample 8 | Sample 8 | [7 empty] | Sample 8
| Sample 9 | Sample 9 | Sample 9 | [7 empty] | Sample 9
| Sample 10 | Sample 10 | Sample 10 | [7 empty] | Sample 10
Same samples - different library preps | RUN 2 - LANE 1 | [9 empty] | RUN 2 - LANE 1
| NA12878 - B - barcode 1 | [9 empty] | NA12878 - B - barcode 1
| NA12878 - B - barcode 2 | [9 empty] | NA12878 - B - barcode 2
| NA12878 - B - barcode 3 | [9 empty] | NA19240 - B
| NA19240 - B | [9 empty] | Sample 4
| Sample 5 | [9 empty] | Sample 5
| Sample 6 | [9 empty] | Sample 6
| Sample 7 | [9 empty] | Sample 7
| Sample 8 | [9 empty] | Sample 8
| Sample 9 | [9 empty] | Sample 9
| Sample 10 | [9 empty] | Sample 10
| RUN 3 - LANE 1 | [9 empty] | RUN 3 - LANE 1
| NA12878 - C - barcode 1 | [9 empty] | NA12878 - C - barcode 1
| NA12878 - C - barcode 2 | [9 empty] | NA12878 - C - barcode 2
| NA12878 - C - barcode 3 | [9 empty] | NA19240 - C
| NA19240 - C | [9 empty] | Sample 4
| Sample 5 | [9 empty] | Sample 5
| Sample 6 | [9 empty] | Sample 6
| Sample 7 | [9 empty] | Sample 7
| Sample 8 | [9 empty] | Sample 8
| Sample 9 | [9 empty] | Sample 9
| Sample 10 | [9 empty] | Sample 10
| ▲ I | [9 empty] | ▲ I
| INTER-RUN Set up 10 samples 3 times from start to finish. This reflects clinical reality of needing to re-run samples | [9 empty] | INTER-RUN Set up 10 samples 3 times from start to finish. This reflects clinical reality of needing to re-run samples

## E_Example_Performance_Metrics
SAMPLE ANALYTICAL PERFORMANCE SUMMARY
The section below provides an example (reflecting a disease-targeted NGS panel of ~250 kb). Only the methods-based validation results (based on NA128778) are shown
GENERAL TECHNICAL PERFORMANCE (across all samples)
Selected Sequencing Metrics
Picard Metrics ("HsMetrics") | Result | Description
Average Total Reads | 129656564.4667 | total number of reads in the bam file
min | 84237058
max | 178976730
Average % Pass-Filter Read Duplication | 7.346% | % reads passing quality filters marked as PCR duplicates
min | 3.385%
max | 23.33%
Average AT Dropout Score | 29.8283 | % reduction in coverage relative to the mean among regions with GC 0-50%
min | 25.7378
max | 43.5624
Coverage Analysis Summary
GATK Metrics ("DepthOfCoverage") | Result | Description
Average Sample Coverage (Raw) | 194.6999 | average depth of coverage excluding reads with no MQ/BQ filtering
min | 123.9142
max | 267.8517
Average Sample Coverage (Filtered) | 193.8019 | average depth of coverage excluding reads with MQ<17 and BQ<10
min | 122.4324
max | 264.6815
Average Sample %Bases >=15 | 99.57% | % bases in exons +/-15 bp with a minimum depth of coverage of 15X (excludes reads with MQ<17 and BQ<10)
min | 99.04%
max | 99.82%
SUMMARY ANALYTICAL PERFORMANCE
Summary of main metrics (details in table below) |  |  |  | Educational example (presence of false positives and false negatives)
Clinical Performance Summary | Version 1 | 95% CI |  | 1 FP, 1 FN | 95% CI | 2 FP, 3 FN | 95% CI
Sensitivity (TP/TP+FN) (All Variants) | 100.00% (145/145) | 97.42-100.00% |  | 99.31% (144/145) | 96.20-99.88% | 97.93% (142/145) | 94.09-99.29%
SNVs | 100.00% (142/142) | 97.37-100.00% |  | 100.00% (142/142) | 97.37-100.00% | 99.30% (141/142) | 96.12-99.88%
In/Dels | 100.00% (3/3) | 43.85-100.00% |  | 66.67% (2/3) | 20.77-93.85% | 33.33% (1/3) | 6.15-79.23%
Specificity (TN/TN+FP) | 100.0000% (0 FP) | 99.9999-100.0000% |  | 99.9996% (1 FP) | 99.9978-99.9999% | 99.9992% (2 FP) | 99.9972-99.9998%
Positive Predictive Value (TP/TP+FP) | 100.00% (0 FP) | 97.42-100.00% |  | 99.31% (1 FP) | 96.20-99.88% | 98.61% (2 FP) | 95.08-99.62%
Intra-Run Variability (Sensitivity Range) | 100.00-100.00% | . |  | 99.31-100.00% | . | 97.93-100.00% | .
Inter-Run Variability (Sensitivity Range) | 100.00-100.00% | . |  | 99.31-100.00% | . | 97.93-100.00% | .
Notes:
The QD/FS filters correctly filter out false positive calls within the region of interest
False Negatives:
none
False Positives:
none
DETAILS - ANALYTICAL PERFORMANCE METRICS
Analytical performance metrics derived from: NIST v2.19 (NA12878) | [7 empty] | Educational example (false positives and false negatives)
| NA12878-1 | NA12878-2A | NA12878-2B | NA12878-3A | NA12878-3B | NA12878-4A | NA12878-4B | 0 FP, 0 FN | 1 FP, 1 FN | 2 FP, 3 FN
Total Bases | 256357 | 256357 | 256357 | 256357 | 256357 | 256357 | 256357 | 256357 | 256357 | 256357 | Total Bases
True Negatives (TN) | 256212 | 256212 | 256212 | 256212 | 256212 | 256212 | 256212 | 256212 | 256211 | 256210 | True Negatives (TN)
False Positives (FP) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 2 | False Positives (FP)
False Negatives (FN) | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 1 | 3 | False Negatives (FN)
True Positives (TP) | 145 | 145 | 145 | 145 | 145 | 145 | 145 | 145 | 144 | 142 | True Positives (TP)
SNVs | 142 | 142 | 142 | 142 | 142 | 142 | 142 | 142 | 142 | 141 | SNVs
In/Dels | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 3 | 2 | 1 | In/Dels
Sensitivity (TP/TP+FN) (All Variants) | 100% | 100% | 100% | 100% | 100% | 100% | 100% | 100% | 99.31% | 97.93% | Sensitivity (TP/TP+FN) (All Variants)
95% Confidence Interval | 97.42-100.00% | 97.42-100.00% | 97.42-100.00% | 97.42-100.00% | 97.42-100.00% | 97.42-100.00% | 97.42-100.00% | 97.42-100.00% | 96.20-99.88% | 94.09-99.29% | 95% Confidence Interval
Sensitivity (SNVs) | 100% | 100% | 100% | 100% | 100% | 100% | 100% | 100% | 100% | 99.3% | Sensitivity (SNVs)
95% Confidence Interval | 97.37-100.00% | 97.37-100.00% | 97.37-100.00% | 97.37-100.00% | 97.37-100.00% | 97.37-100.00% | 97.37-100.00% | 97.37-100.00% | 97.37-100.00% | 96.12-99.88% | 95% Confidence Interval
Sensitivity (In/Dels) | 100% | 100% | 100% | 100% | 100% | 100% | 100% | 100% | 66.67% | 33.33% | Sensitivity (In/Dels)
95% Confidence Interval | 43.85-100.00% | 43.85-100.00% | 43.85-100.00% | 43.85-100.00% | 43.85-100.00% | 43.85-100.00% | 43.85-100.00% | 43.85-100.00% | 20.77-93.85% | 6.15-79.23% | 95% Confidence Interval
Specificity (TN/TN+FP) | 100% | 100% | 100% | 100% | 100% | 100% | 100% | 100% | 100% | 100% | Specificity (TN/TN+FP)
95% Confidence Interval | 99.9999-100.0000% | 99.9999-100.0000% | 99.9999-100.0000% | 99.9999-100.0000% | 99.9999-100.0000% | 99.9999-100.0000% | 99.9999-100.0000% | 99.9999-100.0000% | 99.9978-99.9999% | 99.9972-99.9998% | 95% Confidence Interval
Positive Predictive Value (TP/TP+FP) | 100% | 100% | 100% | 100% | 100% | 100% | 100% | 100% | 99.31% | 98.61% | Positive Predictive Value (TP/TP+FP)
95% Confidence Interval | 97.42-100.00% | 97.42-100.00% | 97.42-100.00% | 97.42-100.00% | 97.42-100.00% | 97.42-100.00% | 97.42-100.00% | 97.42-100.00% | 96.20-99.88% | 95.08-99.62% | 95% Confidence Interval
Notes:
The QD/FS filters correctly filter out false positive calls within the region of interest
False Negatives:
none
False Positives:
none

## F_Assay_Limitations
ASSAY LIMITATIONS - GENERAL GUIDANCE
The following questions are designed to guide the laboratory towards formulating limitations
1. Is the analytical Sensitivity sufficiently high for the gene/s and disease/s of interest?
Are there existing benchmarks (e.g. for well established genes) that must be met?
Example 1: CFTR offered for diagnostic testing: This is a well established gene, must be able to detect all common pathogenic CFTR variants. If the validation result shows suboptimal performance --> need to go back to further optimize performance
Example 2: STRC offered for diagnostic testing: No benchmark exists due to technical difficulties associated with this gene (high homology). In this case, it is sufficient to disclose limitations
2. Is the analytical specificity sufficiently high?
If not 100%, orthogonal confirmation may need to be performed to rule out false positives.
3. Are all genes relevant to the test indication covered?
If not, this must be clearly stated in the methodology